## Request for Quotation
RFQ Date تاريخ طلب العرض |  | 2026-06-02 |  | Quotation Valid Till العرض صالح لغاية
PR Reference رمز طلب الشراء |  |  |  | 26060201
Supplier's Information معلومات المزود
Commercial Name الاسم التجاري |  |  |  | Contact Person شخص التواصل لدى المزود
Address عنوان المزود |  |  |  | Mobile/ موبايل
To be filled by SRD supply chain team حقول يتم تعبئتها من قبل فريق المشتريات | [5 empty] | To be filled by supplier حقول يتم تعبئتها من قبل المزود
Item # تسلسل المادة | DESCRIPTION & ITEMS SPECIFICATION البيان ومواصفات المواد |  | Unit الوحدة | Quantity الكمية | Period المدة | Unit Price سعر الوحدة | Total Price السعر الكامل | Brand & Model Details العلامة التجارية والموديل | Notes ملاحظات
1 | Medical equipments for ( Al Farat PHC - Lower Shyookh PHC- Abu Qalqal PHC) as attached annex 1 أجهزة طبية للمراكز الرعاية الصحية الأولية كما هو مرفق 1 |  | Kit | 3 | 1
2 | Medical lab consumbles for ( Al Farat PHC - Lower Shyookh PHC- Abu Qalqal PHC) as attached annex 2 مواد مخبرية و مستهلكات مخبرية للمراكز الرعاية الصحية الأولية كما هو مرفق2 |  | Kit | 3 | 1
Sub-Total | [7 empty] | المجموع
Period | [6 empty] | 1 | المدة
Sales Tax (if applicable) | [7 empty] | الضريبة (إن وجدت)
Other charges (if applicable) | [7 empty] | مصاريف أخرى (إن وجدت)
Discount (if applicable) | [7 empty] | الخصم على السعر (إن وجد)
Currency العملة |  | USD دولار امريكي | Total Cost |  |  |  |  | السعر الكامل
الليرة السورية
Exchage Rate سعر الصرف |  |  |  | Date of Exchage Rate تاريخ سعر الصرف
Payment Terms شروط الدفع |  | All offers will be issued in US Dollars but the payment will be in US Dollars only through Ziraat Bankası or Turkish Lira at the exchange rate of Ziraat Bankası bank on the invoice date ONLY (we can't accept any offer in any other Currency |  | Lead Time for Delivery المدة المطلوبة للتوصيل
Other Requirements متطلبات أخرى |  | ملاحظة 1 يجب ان تكون الأجهزة المطلوبة مطابقة للوصف و مرفقة بكفالة لا تقل عن 3 سنوات و كتييبات التشغيل و الصيانة و المعايرة و كيفية تنظيف و صيلنة الدورية يجب على من يتقدم للعرض الاجهزة ان يقدم عرض للمواد المخبرية تتناسب و تتوافق مع الاجهزة المقدمة بالعرض و كذلك يجب تشغيل و تجريب الاجهزة في مكان الاستلام Note 1: The requested devices must conform to the stated specifications and be supplied with a warranty of not less than three years, along with operation, maintenance, and calibration manuals, as well as instructions for cleaning and periodic maintenance. Any bidder submitting an offer for the devices must also submit an offer for laboratory materials that are suitable for and compatible with the devices included in the offer. The devices must also be operated and tested at the place of delivery/receipt. |  | Delivery Address عنوان تسليم المواد |  |  | Delivery of the items direct to the health 3centers (Al Farat PHC - Lower Shyookh PHC- Abu Qalqal PHC) Aleppo Governorate.
Offer Received By مستلم العرض
Terms and Conditions الشروط العامة |  | العرض قابل للتجزئة ويحق للمنظمة تثبيت جزء من عرض السعر Offer is Separate-able, SRD has the right to take part of it على المزود التوقيع فوق الختم مباشرة, و ارفاق صورة عن الهوية الشخصية هو شرط لازم لقبول العرض Supplier's Signature should be over companys stamp, and providing copy of ID is a MUST to accept the offer
Supplier's Name, Signature, Stamp & Date اسم المزود و توقيعه وختمه و التاريخ

## Annex 1
PR 26060201 Medical equipments Annex 1
# | Item Description |  | QTY/Facility | Number of facilities | Unit Price | Total | Place of Origin | Model
1 | Stethoscope set | Dual-head (dual face) stethoscope – diaphragm & bell Chestpiece Material Machined stainless steel or aluminum alloy Diaphragm Tunable, high acoustic sensitivity; diameter ~45 mm Bell Small bell (~30 mm diameter) for low-frequency sounds Tubing Latex-free PVC, length ~55–70 cm, good sound insulation Eartips Soft-sealing, replaceable, latex-free | 2 | 1
2 | Otoscope set | Light Source High-intensity LED illumination Power Supply 2 × AA batteries Magnification Integrated optical lens with ≥ 3× magnification Accessories Set of disposable specula, carrying case | 1 | 1
3 | Nebuliser | medication into aerosol mist for respiratory therapy. The device shall operate on 220–240V, 50/60Hz power supply, with a medication cup capacity of approximately 6–8 ml. Particle size shall be within the respirable range, preferably around 0.5–5 microns or MMAD approximately 3–5 microns. Average nebulization rate shall be at least 0.2 ml/min. Compressor flow shall be approximately 4–8 L/min. Noise level shall not exceed 65 dB. The nebulizer shall be supplied with adult mask, pediatric mask, mouthpiece, nebulizer cup, air tube, and air filters. The device shall be compact, easy to use, easy to clean, and suitable for repeated use in primary health care settings. Warranty shall be at least one year. Preferred model: any basic compressor nebulizer or equivalent meeting the above specifications. | 2 | 1
4 | Sphygmomanometer | Manual mercury-free aneroid sphygmomanometer, calibrated, with pressure range 0–300 mmHg and accuracy ±3 mmHg, supplied with adult nylon cuff with Velcro fastening, latex-free bladder and tubing, inflation bulb, precision air-release valve, single-head stethoscope, and zippered carrying pouch. | 3 | 1
5 | Oxygen cylender with flowmeter | Medical oxygen cylinder with flowmeter, seamless steel or aluminum, 40-liter water capacity, working pressure 150–200 bar or 2000 PSI, equipped with manual cylinder valve, oxygen pressure regulator, pressure gauge, adjustable flowmeter 0–15 L/min, humidifier bottle, valve wrench, and oxygen connection tube. The cylinder shall be suitable for medical oxygen use and supplied with valid safety/hydrostatic test certificate. | 1 | 1
6 | Serum holder | IV Stand with Four Hooks Four strong hooks for safely hanging medical IV fluids. Sturdy base with wheels to ensure stability and easy movement. Adjustable height to suit all medical uses. Practical and lightweight design for easy transport. Rust-resistant and easy to disinfect for repeated use in medical environments. | 1 | 1
7 | Puls oximeter | neonatal wrap probe and adult reusable probe.Portable pulse oximeter device designed for continuous or spot-check monitoring of oxygen saturation (SpO₂) and pulse rate in patients of all ages, including neonates and adults, suitable for use in Basic Emergency Obstetric and Newborn Care (BEmONC) settings SpO2 range: 0-100% (±2% accuracy); Pulse rate: 25-250 BPM (±3 BPM); Probes: Reusable neonatal/adult finger clip; Display: LCD/OLED; | 3 | 1
8 | Trolly | طاولة معدنية ستانليس ستيل تتكون من درجين و رف سفلي بالاضفة الى 4 عجلات مع حركة دورانبة و اقفال للعجلات بقياس يتجاوز 90 سم A stainless steel table comprising two drawers and a lower shelf, as well as four swivel castors with locks, measuring over 90 cm | 2 | 1
9 | Echography device | Applications: Abdomen, OB/GYN, Vascular, Small parts (breast, testicles, thyroid), Musculoskeletal Conventional & Superficial, Pediatric, Basic Cardiac Imaging technologies: i-Image, SRA, Compound imaging, THI Display Modes: B, B/B, 4B, M, B/M, PW Gray Scale: 256 Levels Scanning Depth: 240mm Max (depend on the probe types) Probe: 3.5 MHz – 5.0 MHz Convex Cine Loop: 256 frames Zoom: 4 steps Software Packages: Dedicated measurement package: General, OB/GYN, URO, Cardiac, Vascular, Small parts Archives Management: Patient information management, images record, exam report Standard Configuration: 1 probe connector, 3 USB ports, VGA port, Video port, LAN port, Full Screen Display Mode, i-Image, Chroma, SRA, Compound imaging, THI, 8G memory card, PW Mode Available: Waterproof keyboard cover, Support PC Printer, Preferred CHISON brand or equivalent | 1 | 1
10 | ECG device | Portable 12-lead ECG machine with 12-channel simultaneous acquisition, built-in thermal printer, LCD display, rechargeable battery, automatic/manual recording modes, heart rate measurement, ECG waveform display, and standard interpretation/measurement function. The device shall include complete lead cable, limb electrodes, chest electrodes, ECG thermal paper, battery, charger/power cable, and user manual, with at least one-year warranty | 1 | 1
11 | glucometer with 500 strip | Sample Size: 0.4 µL (microliter) of blood. Test time: 5-10 sec; Memory: 500 results; Range: 20-600 mg/dL; Power Source: lithium coin battery | 1 | 1
12 | Adult scale | Capacity 0–200 kg (minimum), preferably up to 250 kg Graduation (Accuracy) ±100 g (0.1 kg) Platform Size Minimum 30 × 30 cm, non-slip surface Display Digital LCD, backlit, easy to read Units kg Power Supply 2 × AAA batteries or rechargeable battery Auto Zero & Auto Off Yes | 1 | 1
13 | Children scale | Capacity 0–20 kg (minimum), preferably up to 25 kg Graduation (Accuracy) ±5 g for weights <10 kg; ±10 g for 10–20 kg Platform Design Curved, comfortable cradle or baby tray with raised sides for safety Display Large digital LCD, backlit for easy reading Units kg Power Supply 2 × AA/AAA batteries or rechargeable battery Auto Zero & Auto Off Yes Tare Function Yes, to subtract blanket or diaper weight Hold Function Yes, to lock reading even if baby moves Material Easy-to-clean, smooth, non-toxic plastic | 2 | 1
14 | Height and weight Schale | Height measure scale-Hall mounted height rods or height rods that mounted into a scale (2M)Digital weight scale in kilograms up to 180 KG | 2 | 1
15 | Examination light | Light intensity: 20,000-50,000 lux; LED (50,000 hours life); Adjustable arm: 80-120 cm; Castors: 50 mm | 3 | 1
16 | Automated hematology analyser | Automated 3-part differential hematology analyser, with throughput of at least 72 samples per hour, using a 2-reagent system including diluent and lyse. The analyser shall provide at least 22 parameters and 3 histograms, including WBC, LYM#, MID#, GRAN#, LYM%, MID%, GRAN%, RBC, HGB, HCT, MCV, MCH, MCHC, RDW-CV, RDW-SD, PLT, MPV, PDW-SD, PDW-CV, P-LCR, P-LCC, and PCT. Principles of operation shall include electrical impedance for WBC/RBC/PLT counting and photoelectric colorimetry for HGB measurement. The device shall include automatic internal and external cleaning for the aspiration probe, internal preheating for diluent, automatic clog removal by aperture cauterization and flushing, and internal storage for lyse. The analyser shall have a high-resolution color LCD screen, approximately 10.4 inches, and a built-in thermal printer. Sample volume shall be low, preferably 8.5 µL or less for whole blood mode and approximately 20 µL for pre-dilute mode. The device shall support LIS and HIS connectivity with HL7 protocol, and shall include RS232, Ethernet network port, and USB ports allowing connection to mouse, keyboard, and external printer. Data storage shall be at least 100,000 sample results with histograms, preferably up to 1,000,000 results. The analyser shall support L-J, X, X-R, and X-B quality control modes, with manual and automatic calibration. The system shall be open or non-locked for reagents, allowing the use of validated compatible reagents without affecting calibration, quality control, or warranty. The supplier shall provide the required operating software/license, installation, calibration, user training, user manual, and all necessary accessories for operation. Preferred model: Spincell 3n or equivalent meeting or exceeding the above specifications | 1 | 1
17 | Spectrophotometer device | Semi-automated biochemistry analyzer / spectrophotometer capable of analyzing clinical biochemistry and turbidimetric assays. The analyzer shall have LED-based maintenance-free optical technology with a spectral range of at least 340–670 nm and photometric range up to 3.5 Abs. The system shall support monochromatic and bichromatic assays, including absorbance, end-point, kinetic, fixed-time, and differential reading modes, with or without reagent blank. The analyzer shall include at least 8 wavelengths/filters covering 340, 405, 505, 535/546, 560/578, 600/620, 635, and 670 nm, with at least one optional/free filter position if applicable. The device shall be able to read through quartz flow cuvette, macro, semi-micro and micro cuvettes, and round test tubes approximately 12 mm x 75 mm. Flow cell volume shall be approximately 18 µL. The analyzer shall include temperature control/incubation system, preferably Peltier-based, covering 25–40°C or at least 25°C, 30°C, and 37°C. The device shall have an LCD or color touchscreen display, built-in thermal printer or printer connectivity, USB port, and Ethernet/PC connectivity where available. The analyzer shall include programmable test methods, quality control management with Levey-Jennings charts, calibration by factor, single calibrator, or multipoint calibration curve, and sufficient internal data storage for patient results, blanks, calibration, and QC data. The system shall support validated compatible reagents and customized test programming where applicable. The supplier shall provide operating software/license, installation, calibration, user training, user manual, power cable, waste bottle, printer paper, and all necessary accessories for operation. Preferred model: BioSystems BTS or equivalent meeting or exceeding the above specifications. | 1 | 1
18 | Microscope | Binocular compound light microscope with extra wide-field eyepieces EW10×/22, reversed quintuple nosepiece, achromatic objectives 4×, 10×, 40×, and 100× oil immersion, providing total magnification of 40×, 100×, 400×, and 1000×. The microscope shall be equipped with a swing-out condenser NA 0.9/0.25, double-layer mechanical stage approximately 185×142 mm with travel range approximately 75×55 mm, coaxial coarse and fine focusing system, built-in transmitted illumination with LED light source or halogen lamp 6V/30W, aspherical collector, and blue and green filters. The microscope shall be supplied with immersion oil, dust cover, power cable, user manual, and all necessary accessories for routine laboratory use. مجهر ضوئي مركب ثنائي العدسة العينية، مزود بعدسات عينية واسعة المجال EW10×/22، وحامل عدسات شيئية خماسي معكوس، وعدسات شيئية لا لونية 4× و10× و40× و100× باستخدام الزيت، بما يوفر تكبيراً كلياً 40× و100× و400× و1000×. يجب أن يكون المجهر مزوداً بمكثف متحرك/قابل للإزاحة NA 0.9/0.25، ومنضدة ميكانيكية مزدوجة الطبقة بقياس تقريبي 185×142 مم وبمدى حركة تقريبي 75×55 مم، ونظام ضبط خشن ودقيق متحد المحور، وإضاءة نافذة مدمجة بمصدر LED أو مصباح هالوجين 6V/30W مع مجمّع لا كروي، بالإضافة إلى فلتر أزرق وفلتر أخضر. يجب أن يزوّد المجهر بزيت الغمر، وغطاء حماية من الغبار، وكابل كهرباء، ودليل استخدام، وجميع الملحقات اللازمة للاستخدام المخبري الروتيني. | 1 | 1
19 | Centerfuge | Digital laboratory centrifuge, bench-top type, with speed range 500–5000 rpm, capacity 12 × 15 ml tubes, digital control panel/display for speed and time adjustment, adjustable timer, brushless motor, low-noise operation, safety lid lock, imbalance protection, and automatic shutdown when the lid is opened. The device shall be supplied with compatible rotor, tube holders/adapters, power cable, user manual, and minimum one-year warrant | 1 | 1
20 | Hematocrit Cinterfuge | Hematocrit Centrifuge: Hematocrit centrifuge with a rotation speed of 10,000–14,000 rpm and relative centrifugal force of 13,000–18,000 ×g. The device has a capacity of 24 capillary tubes and is supplied with a numbered hematocrit reading disc. It is equipped with a safety cover/lid for secure operation, a digital display to show speed and time, and a brushless DC motor to ensure quiet and stable operation. العربية: جهاز طرد مركزي للهيماتوكريت: جهاز طرد مركزي للهيماتوكريت بسرعة دوران تتراوح بين 10,000–14,000 دورة في الدقيقة، وبقوة طرد مركزي نسبية تتراوح بين 13,000–18,000 ×g. تبلغ سعة الجهاز 24 أنبوباً شعرياً، ويزوّد مع قرص قراءة هيماتوكريت مرقّم. كما يكون مزوداً بغطاء أمان لضمان التشغيل الآمن، وشاشة رقمية لعرض السرعة والوقت، ومحرك تيار مستمر عديم الفرشات لضمان تشغيل هادئ ومستقر. | 1 | 1
21 | Water bath | Type: Digital controlled laboratory water bath Temperature Range: Ambient +5 °C to 100 °C (±0.5 °C accuracy) Temperature Control: Thermostatic control with digital display Capacity: 6–12 liters (common size for PHC labs) Heating Element: Immersion-type stainless steel heater Material (Tank): Corrosion-resistant stainless steel Material (Body): Powder-coated steel or epoxy-coated steel housing Lid: Hinged or removable stainless-steel lid to reduce evaporation Safety Features: Over-temperature cut-off, low-water level protection Accessories: Test tube racks, flask holders Compliance: equivalent certification for medical/lab use | 1 | 1
22 | Tube Mixer | Tube Mixer: speed range, tube capacity, timer, vortex/roller typ | 1 | 1
23 | Micropipette set | Photometric range: 0.5–10 µL, 10–100 µL, 100–1000 µL. Abs for all wavelengths. | 1 | 1
24 | Stop watch | digital stopwatch, start/stop/reset, lap time, battery. | 1 | 1
25 | Tempreture and humidity device | temperature range, humidity range, accuracy, display, battery. | 3 | 1
26 | ESR set (stand and tube) | ESR stand, Westergren tubes, rack, pipettes, timer. | 1 | 1
27 | Examination table with steps | مصنوع من الفولاذ القوي المطلي بالبودرة (Powder-coated) المقاوم للصدأ والتآكل .السطح (المرتبة): حشوة إسفنجية عالية الكثافة مغطاة بجلد "فينيل" طبي، مقاوم للبكتيريا، العفن، اللهب، وسهل التنظيف والتطهير. التعديل: إمكانية تعديل زاوية مسند الظهر (من 0 إلى 55 أو 80 درجة) لتوفير وضعية جلوس مريحةللمريض .الارتفاع: ارتفاع ثابت (قياسي) الإضافات درج صعود معدني او خشبي ارتفاع درجة الواحدة ١٥ سم .سعة الوزن: مصممة لتحمل أوزان ثقيلة، تبدأ من ٥٠ كجم وتصل إلى أكثر من 220 كجم | 4 | 1
28 | Privacy screen/Paravan | 3-panel folding screen, washable curtain, metal frame, wheels. | 4 | 1
29 | Electrical Suction Machine | suction pressure, jar capacity, flow rate, foot switch/tubing/filter. | 1 | 1
30 | Minor Basic Instrument stanlesstil Set with metal box | Set contains as follows Blade Holder, 22 mm – 1 piece Blade Holder, 15 mm – 1 piece Needle Holder (Port Quin) – 1 piece Forceps with Tooth – 1 piece Forceps without Tooth – 1 piece Regular Scissors – 1 piece Mayo Scissors – 1 piece Kocher Clamp – 1 piece Kelly Clamp – 3 pieces Decontamination Forceps – 1 piece Tool Basket – 1 piece | 2 | 1
31 | Dry sterilizer | dry heat sterilizer, chamber size, temperature range, timer, stainless steel chamber. | 1 | 1
32 | Blodd collection chairPhlebotomy Chair | Comfortable blood collection chair suitable for phlebotomy and sample collection procedures, with a strong metal frame and easy-to-clean leatherette or medical-grade PVC upholstery. The chair shall be equipped with padded armrests, including at least one adjustable armrest suitable for blood drawing, with the possibility of right or left arm use. The chair shall have a stable base, anti-slip feet, and a comfortable backrest and seat. Upholstery shall be waterproof, disinfectant-resistant, and easy to clean. The chair shall support a minimum patient weight of 120–150 kg. The design shall be suitable for repeated use in laboratory and primary health care settings | 1 | 1
Total Cost

## Annex 2
PR 26060201 Medical lab consumbles Annex 2
ID | Description | Unit | Quantity | UnitCost | TotalCost
21 | رؤوس ممصات زرقاء Blue micropipette tips | box 500 pcs | 8
22 | رؤوس ممصات صفر yellow micropipette tips | box 1000 pcs | 8
23 | high quality Urine strips(Germany ) شرائط بول كمبور عالية الجودة (الماني) | box 150 pcs | 10
26 | سواتر عادية coverslip size 22 mm × 22 mm | big box contain 10 small boxes | 2
27 | cornial bottom plastic tubes for urine أنابيب مخروطية القعر بلاستيكية لتحليل البول | box 250 pcs | 8
28 | Tubes for chemical tests أنابيب بلاستيكية 5 مل لإجراء تحاليل الكيميائية | box 250 pcs | 15
29 | أنابيب pt سيترات الصوديوم 13x75 mm blood collection tubes with Sodium Citrat for PT tests | box 100 pcs | 2
30 | أنابيب كيميا هيبارين الليثيوم 13x75 mm blood collection tubes with Lithium Heparin for biochemical tests | box 100 pcs | 6
31 | أنابيب دمويات تعداد 13x75 mm blood collection tubes with EDTA for CBC tests(K2-K3) | box 100 pcs | 5
32 | CRP quantitative (BioSystems or equivalent) reagent A 2 x 160 mL, Reagent 2 x 40 mL with standerd سي آر بي كمي( بيوسيستم أو مايعادله) | kit | 2
33 | kit Heamatology analyzer( ) Diluent 20 liter , layz 1 liter . | kit | 4
34 | Creatinin reagent (BioSystems or equivalent)A 2 x 50 mL, B2 x 50 Ml with standerd كاشف كرياتينين ( بيوسيستم أو مايعادله) | kit | 4
35 | كاشف حمض البول لونيuric acid reagent A*200ml ( biosystem or equivalent ) | BOX | 2
36 | كاشف شحوم ثلاثية triglycerid reagent A4*50ml | BOX | 2
37 | كاشف كوليسترول لوني cholesterol reagent A*200ml (biosystem or equivalent ) | BOX | 2
38 | Capillary tubes for hematocrit with Heparin أنابيب شعرية مهبرنة للهيماتوكريت | Box (10 cans) صندوق 10 علب | 2
39 | Blood groups set high quality (Anti A,Anti B,Anti RH) A,B,RH | SET | 10
40 | بلاكات اختبار حمل (آكون ACON أو ما يعادله ) Pregnency test strip kit Beta HCG | BOX 40 TEST | 20
41 | كاشف ألبومين ( بيوسيستم أو مايعادله) Albumin 2x250 ml | set | 1
42 | Billirubine 200ml with st kit spinreact | set | 2
43 | ASLO quantitative (BioSystems or equivalent) reagent A 1 x 40 mL, Reagent 1 x 10 mL with standerd | BOX | 1
44 | كاشف بولة ( بيوسيستم أو مايعادله) لوني A 1x 250 mL, B 1 x 250 mL Urea reagent (BioSystems or equivalent) | BOX | 2
45 | Widal Test 10ML (O,H)american | Btl | 2
46 | Wright test 10MLamerican | Btl | 2
47 | S GPT Test Kit 200 ml | Kit | 2
48 | كاشف سكر ( بيوسيستم أو مايعادله) 1X1000 ML Glucose reagent(BioSystems of equivalent) | kit | 4
49 | كاشف بروتين بول ( بيوسيستم أو ما يعادله) Protein, urine or CSF 4 x 50 mL reagent A | kit | 2
50 | ESR tube,Fl, Black | BOX 100PCS | 1
51 | شرائح زجاجية Glass slides for microscope (box of 100) | box 100 PCS | 40
52 | platelet reagent محلول صفيحات سيربيو أو ما يعادله | bottle 100 ml | 2
53 | CBC printer roll 5.5 cm | pcs | 10
54 | Concentrated Asetic Acid حمض الخل الثلجي المركز | bottle 100 ml | 2
Total